## Slide 1: 7.15 group meeting
Pose calibration and force mapping
7 experiment
6628 matched motion pairs
Manual script update
2428 samples
8D proxy input to normal force
Policy models
2 baselines
ACT and diffusion training pipelines
01
Ultrasound probe imitation learning data collection and model training

## Slide 2
This week turned calibration into manuscript-ready evidence
1
Rigid pose transformation
Validated tag-to-flange conversion on full and axis-constrained robot motions.
2
Normal force mapping
Trained proxy-signal mapping from IMU and Paxini measurements to KWR75B normal force.
3
Manuscript integration
Added result figure, table, and polished Section VI-A text for rigid transformation validation.
4
Imitation learning baselines
Built ACT-style and diffusion-policy-style training code; model training is running.
Main takeaway
The data migration pipeline now has quantitative calibration evidence, while the learning baselines are in place for the next group-meeting update.
Ultrasound probe imitation learning - weekly update
02

## Slide 3
Force mapping now has a trained calibration model
Portable proxy signals are mapped to robot-frame normal force.
Dataset
2428
aligned samples from 7 sessions
Input
8D
pitch, roll, left/right 3-axis Paxini force
Target
1D
KWR75B normal force
Split
1699 / 364 / 365
train / val / test
Model performance
Split
MAE (N)
RMSE (N)
R2
Pearson
Train
0.826
1.096
0.910
0.955
Validation
1.067
1.505
0.860
0.930
Test
1.183
1.683
0.847
0.921
Held-out interpretation
- Test MAE: 1.18 N; test RMSE: 1.68 N.
- Test R2: 0.847; Pearson correlation: 0.921.
- Validation and test gaps are moderate, suggesting the mapping generalizes beyond training sessions.
Ultrasound probe imitation learning - weekly update
03

## Slide 4
Rigid pose conversion: 6628 matched motion pairs
Full 6-DoF test error
2.73 mm
translation mean
Full 6-DoF rotation
1.24 deg
rotation mean
Pairing quality
16.72 ms
maximum test-set time difference
Ultrasound probe imitation learning - weekly update
04

## Slide 5
The manuscript now has quantitative support for Section VI-A
Added artifacts
- Rigid transformation validation figure in PNG/PDF/SVG.
- LaTeX table summarizing seven calibration sequences.
- Polished manuscript paragraph for Section VI-A.
- Large-font figure version for readability in manuscript and slides.
What the section now argues
The calibrated marker-to-flange transform is stable on held-out motion pairs and can be used as a fixed conversion parameter for moving hospital-side marker trajectories into the robot end-effector frame.
Ultrasound probe imitation learning - weekly update
05

## Slide 6
ACT and diffusion baselines are ready for training results
ACT-style chunk regressor
ultrasound image -> ResNet18 -> MLP -> future pose delta chunk
- Action horizon defaults to 20 x 7 pose deltas.
- Supports with-force and without-force dataset variants.
- Splits by source session to reduce temporal leakage.
Diffusion-policy-style model
noisy future chunk + timestep + image condition -> MLP denoiser
- Uses ResNet18 image condition.
- Force can be concatenated to the conditioning vector.
- Training scripts and configs are in place.
Ultrasound probe imitation learning - weekly update
06

## Slide 7
Training results remain the main group-meeting update
Data conversion and calibration evidence
Manuscript Section VI-A update
ACT training
Diffusion training
Placeholder for group meeting
Insert ACT/diffusion loss curves and validation metrics once the current runs finish.
Ultrasound probe imitation learning - weekly update
07

## Slide 8
Next step: connect calibrated data to policy learning
Before group meeting
- Add ACT and diffusion loss curves.
- Report validation and test metrics.
- Identify whether force input improves pose prediction.
Manuscript updates
- Place figure between related result paragraphs.
- Keep large-font figure for readability.
- Align table and caption with final results wording.
Experiment follow-up
- Check robustness across more scanning sessions.
- Compare with-force vs without-force policies.
- Prepare qualitative rollout examples if available.
Goal for the next update: move from calibrated data conversion to measurable imitation-learning performance.
Ultrasound probe imitation learning - weekly update
08